Checklist for TPC after long shutdown:                          valid 7/14/2009

All tasks to be performed BEFORE pole tips are inserted.

# CABLES

1. Check all cable connections on TPC face that were

uncabled for maintenance work:

Check labels						____

Check ground sleeves					____

# ANODES

1. Both VME processors boot and ARCNET OK			____

2. Both serial sessions OK						____

3. Ramp HV to 100 V (N2 in chamber)				____

4. All channels draw ~500 nA on ramp up?				____

5. No channels trip							____

6. Check DC current at 100 V – any new high currents?		____

7. Check hardware trip limit for all cards:

Inner = 1210 V					____

Outer = 1500 V					____

8. Test interlock from PLC

# CATHODE & FIELD CAGE

1. Field cage current and voltage read out by Kiethley?		____

2. Remote power switch turns Glassman on/off?			____

3. Glassman turns on remotely? Slow controls)			____

4. Ramp Glassman to 1000V (N2 in chamber)			____

5. Check field cage currents and voltages.				____

6. Test interlock from PLC

# GATED GRID

1. Both crates (control & driver) OK?				____

2. Check capacitance for each cable/sector – pin to pin & pins to

ground								____

2A. Check cable ground braid to platform ground			____

3. Turn on GG – download setpoints					____

4. Pulse GG with pulser – check monitor out for all sectors		____

5. Calibrate outputs if needed.					____

6. Check that ground sleeves in back of rack are in place.		____

7. When done, ensure trigger cable from TCD is plugged in		____

8. Test interlock from PLC						____

# GROUND PLANE PULSER

1. Wavetek on and downloaded pulse selected?			____

2. Downloaded pulse looks ok? (Parameters stored correctly?)	____

3. Rate limiter plugged in?						____

4. Check that rate limiter is working before plugging into fanouts	____

5. Trigger Wavetek with pulser – check all outputs from fanouts	____

6. Replace bad fanout modules.                                                         ____

7. Make sure trigger cable from TCD plugged into Wavetek		____

8. Check PLC interlock                                                                       ____

# FEES & MWC FEES

1. Check that all blowers in rack row 2B are running		____

2. Turn on VME crate and boot processor				____

3. Make sure all manual switches on LVPS are in remote position	____

4. Turn on all FEEs and MWC FEEs	remotely.			____

5. Turn off TPC water skid – do all FEEs go off?			____

6. TPCTEMP computer running and updating?			____

# LASER

1. Power on both lasers – locally					____

2. Beams aligned?							____

3. Photodiode signal ok to TCD and cathode TDC?			____

4. Check remote on/off of lasers (slow controls)			____

5. Check remote viewing of lasers					____

6. Test PLC interlock							____

# TPC & GLOBAL INTERLOCKS

1. Check inner field cage air blower – on & flow switch ok?	____

2. Wet, in turn, east and west TPC Tracetek – alarm & skid stops?	____

3. Platform methane sniffer ok? Contact CAS if not			____

# GAS SYSTEM

1. Close SV18 and do leak rate test with N2				____

2. Calibrate and test trips of Pioneer methane system		____

3. Calibrate and test trip of gap methane detector			____

4. Calibrate CAI methane analyzers					____

5. Check limits and test hardware & PC alarms.			____

6. Bake out dryer & purifier						____

7. Check two big compressors					____

8. Test all gas pipe clamps and tighten				____

# ELECTRONICS

1. Do a geometry run – compare to baseline				____

2.Do a pulser run – compare to baseline				____

3. Do a full pedestal run – check RMS offline		            ____