- Attachment 3. STAR TPC GAS SYSTEM CHECK LIST           Valid from 11/1/2007

(To be filled out once per shift and placed in Gas System Binder)

- Computer Monitored Sensors

| Sensor | Function | Value | Range | Comments |
| --- | --- | --- | --- | --- |
| PI1 (PT1) | Output of Big Compressor |  | 95 - 125 mbar |  |
| PT2 | Input to TPC |  | 2.0 – 4.5 mbar |  |
| PT3 | Monitor Chamber Pressure |  | 0.3 – 2.0 mbar |  |
| PT5 | Input of Big Compressor |  | 0.5 – 1.6 mbar |  |
| PT7 | TPC/Gap Differential Pressure |  | 0 – 2.0 mbar |  |
| PT8 | Input Pressure to TPC |  | 1.5 - 2.5 mbar |  |
| PT9 | Pressure across Dryer Filter |  | 10 - 90 mbar | < 0 If dryer loop off |
| PT10 | Pressure across Main Filter |  | 2.0 - 20 mbar |  |
| FT1 | Exhaust Rate |  | -6.0 – +20.0 lpm | Oscillates |
| PTB | New barometer |  | 970 – 1040 mbar |  |
| PI8 | Argon Delivery Pressure |  | 1.0 - 1.7 bar |  |
| PI9 | Methane Delivery Pressure |  | 0.98 - 1.5 bar |  |
| PI10 | Nitrogen Pressure on Pad |  | 2.2 - 4.0 bar |  |
| PI13 | Argon Pressure on Pad |  | 2.2 – 4.3 bar |  |
| PI14 | Methane Pressure on Pad |  | 5 - 155 bar |  |
| PI15 | Nitrogen Delivery Pressure |  | 0.9 - 1.8 bar |  |
| FM1 | Methane mass flow controller |  | 0.9 - 1.9 lpm |  |
| FM5 | Argon mass flow controller |  | 10 – 20 lpm |  |
| FI7 | Recirculating flow in TPC |  | 450 - 600 lpm |  |
| O2 M1 | Oxygen content |  | 5 - 80 ppm |  |
| O2 M5 | Oxygen content |  | < 0 % |  |
| H2O M2 | Water Content |  | 0 - 80 ppm |  |
| CH4 M3(a) | Methane Content |  | 9.5 - 10.5 % |  |
| CH4 M4 | Methane Content |  | 9.5 - 10.5 % |  |
| Leakage | Calculated Leak Rate |  | 0 – 14.5 lpm |  |
| CH4FM% | Methane/Argon Ratio |  | 9.0 – 11.0 |  |

Rack 1

| FM5 | Fresh Argon Flow |  | 10 - 20 lpm |  |
| --- | --- | --- | --- | --- |
| FM1 | Fresh Methane Flow (slaved) |  | 1.0 - 1.9 lpm |  |
| FI7 | Recirculating Flow to TPC |  | 500 - 600 lpm |  |
| PT8 | TPC Pressure |  | 1200 - 2600 bar |  |
| M2 | Water Content |  | 0 - 80 ppm |  |

TURN OVER FOR MORE!

- Rack 2

| Sensor | Function | Value | Range | Comments |
| --- | --- | --- | --- | --- |
| PI7 | Input to Big Compressor |  | 6.0 - 12. mm H2O |  |
| PI2 | Output of SC#1 |  | 8.2 – 12.5 PSIA |  |
| PI1 | Output of Big Compressor 1 |  | 40 - 48 in H2O | Oscillates |
| PI3 | Intermediate Input Pressure |  | 5.0 – 12.0 in H2O |  |
| PI4 | Input To TPC |  | 1.0 - 1.55 in H2O |  |
| FI1a | Gas to Vent |  | 0 – 50 SCFH | Oscillates |
| M4 | Fresh Methane Content |  | 9.5-10.5% |  |
| FI2a | M4 P10 flow |  | 4.0-6.0 x 100 cc/min |  |
| SV14 | Input to M1, M2, M3a |  | Open (Green) |  |
| SV6,7,8 | Input to M1,M2,M3a |  | Open (Green) |  |
| TIC1 | Purifier Temperature |  | 205 - 240 |  |
| PI8a | Output of SC#2 (Back) |  | 0.4 – 0.8 Bar |  |
| FI3 | M2 Flow (Back) |  | 9-12x100 cc/min |  |

Rack 3

| Sensor | Function | Value | Range | Comments |
| --- | --- | --- | --- | --- |
| M1 | Oxygen Meter |  | 5 - 80 ppm |  |
| M5 | Oxygen Meter |  | < 80 ppm |  |
| PI8 | Argon Delivery Pressure |  | 15.0 – 19.0 PSIG |  |
| PI9 | Methane Delivery Pressure |  | 15 PSIG +/- 1 | Must be less than PI8 |
| PI15 | Nitrogen Delivery Pressure |  | 13 - 17 PSIG |  |
| M3b | Methane Analyzer |  | 9.5 - 10.5 % |  |
| SV21 | Methane Inlet |  | Open (Green) |  |
| SV22 | Argon-Methane Bypass |  | Closed (Red) |  |
| FI2 | M1 Flow |  | 0.1-0.3 lpm |  |
| FI4 | M3a Flow |  | 4.0 -6.0 x100 cc/min |  |
| FI10 | Insulation Gap Flow |  | 10.0 – 20.0 lpm |  |
| FI12 | M4 Case Purge (Back) |  | 2.0 – 4.0 lpm |  |
| FI13 | M3a Case Purge (Back) |  | 2.0 - 4.0 lpm |  |

Mixing Room

| Sensor | Function | Value | Range | Comments |
| --- | --- | --- | --- | --- |
| PI10 | Nitrogen Supply Pressure |  | 26 - 52 PSIG |  |
| PI11 | Nitrogen Delivery Pressure |  | 15 - 22 PSIG |  |
| PI12 | Argon Delivery Pressure |  | 15 - 25 PSIG |  |
| PI13 | Argon Supply Pressure |  | 30 - 45 PSIG |  |
| FI16 | Laser Nitrogen Flow |  | < 10 lpm |  |
| FI17 | Water Skid Nitrogen Flow |  | 1- - 20 lpm |  |
| PVENT | Vent Line Pressure |  | 0 – 0.5 in H2O |  |
| ODH | Room Oxygen content |  | 19 – 22 % |  |
| Bubbler | Oil Level |  | Between Marks |  |

Rack 4

| M6 | Gap Oxygen |  | 0 – 40 ppm |  |
| --- | --- | --- | --- | --- |
| M7 | Gap Water |  | 0 – 40 ppm |  |
| M8 | Gap Methane |  | 0 – 16 %LEL |  |
| FI51 | Gap N2 flow |  | 5 – 20 lpm |  |
| FI55 | Gap return flow |  | 0.8 – 3 lpm |  |

Software Alarms Enabled? ___________

Chilled water flow to gas system (gpm) ________

Operator__________________________		Date & Time____________