2/19/2008

TPC tasks from Run 8 shutdown to start of Run 9

Shutdown after Run 8:

- Purge TPC with N2, sign the blue sheet when purged

- Shutdown gas system and establish summer maintenance N2 flow.

- Monitor LAr tank and order refill when needed

- If STAR doesn’t roll, nothing else to do for gas system.

- If STAR rolls, make sure TPC gets N2 during roll and in assembly building.

- When (if) STAR rolls back, put up gas pipes to wall.

- Shutdown platform crates and breakers - everything off except interlock system.

- Remove pulser and put in office (plug it in!)

TPX work over summer:

- Contact Danny and Bob Sheetz - offer assistance for FEE testing etc

- Supervise removal of old electronics - help in “GO” decision for removal.

- Supervise installation of new FEEs, RDOs and grounding cards.

- Supervise (help) installation of new cables - new fiber optics from platform to TPC face and from platform to DAQ room.

- Supervise (help) installation of new LVPS (138 new power supplies in existing boxes) Test slow controls of new LVPS and interlocks.

- Supervise (help) installation of new DAQ PCs in DAQ room.

- As new electronics is installed on TPC start testing with pulser and pad monitor.

- Do R&D on new Alice prototype gated grid driver - initiate construction if needed.

Other summer projects:

- Work with Creighton on slow controls migration from Solaris to Linux. Also possible replacement of  some VME processors with PC based controls - test any new software.

- If work is done in the IFC, make sure to test for shorts.

RUN 9 startup:

- When Leonid comes in the fall, turn on gas system racks and PC.

- Calibrate methane analyzers, pressure transducers, reconstitute purifier and dryer, test system leak rate.

- Calibrate and test interlocks for gas room Pioneer methane system (3 sensors) and gap methane sensor.

- Make sure Bob calls for maintenance on the gas room 3 phase UPS.

- Work with CAD during certification of STAR global interlock system - confirm actions sent to TPC interlock system.

- Turn on all platform racks and crates - reinstall pulser.

- Perform TPC checklist - see separate sheets.

- Confirm with Alexei the laser status - is a site visit by SpectraPhysics needed?

- Perform pulser run - check with pad monitor. Give Danny list of bad channels for replacement.

RUN 9 Physics:

- Three days before “start of physics”, start Ar purge of TPC (typically 6 volumes = 2 days).

- If blue sheet is signed and shift crews are in place, start P10 purge of TPC - 2 volumes.

- Start gas system recirculation - all alarms active, all interlocks green.

- Turn HV on chamber and make laser run - check the position of the membrane flash to check drift velocity.

Good Luck!

Checklist for TPC after long shutdown:

All tasks to be performed BEFORE pole tips are inserted.

# CABLES

1. Check all cable connections on TPC face that were

uncabled for maintenance work:

Check labels						____

Check ground sleeves					____

# ANODES

1. Both VME processors boot and ARCNET OK			____

2. Both serial sessions OK						____

3. Ramp HV to 100 V (N2 in chamber)				____

4. All channels draw ~500 nA on ramp up?				____

5. No channels trip							____

6. Check DC current at 100 V – any new high currents?		____

7. Check hardware trip limit for all cards:

Inner = 1210 V					____

Outer = 1500 V					____

8. Test interlock from PLC

# CATHODE & FIELD CAGE

1. Field cage current and voltage read out by Kiethley?		____

2. Remote power switch turns Glassman on/off?			____

3. Glassman turns on remotely? Slow controls)			____

4. Ramp Glassman to 1000V (N2 in chamber)			____

5. Check field cage currents and voltages.				____

6. Test interlock from PLC

# GATED GRID

1. Both crates (control & driver) OK?				____

2. Check capacitance for each cable/sector – pin to pin & pins to

ground								____

2A. Check cable ground braid to platform ground			____

3. Turn on GG – download setpoints					____

4. Pulse GG with pulser – check monitor out for all sectors		____

5. Calibrate outputs if needed.					____

6. Check that ground sleeves in back of rack are in place.		____

7. When done, ensure trigger cable from TCD is plugged in		____

8. Test interlock from PLC						____

# GROUND PLANE PULSER

1. Wavetek on and downloaded pulse selected?			____

2. Downloaded pulse looks ok? (Parameters stored correctly?)	____

3. Rate limiter plugged in?						____

4. Check that rate limiter is working before plugging into fanouts	____

5. Trigger Wavetek with pulser – check all outputs from fanouts	____

6. Replace bad fanout modules.                                                         ____

7. Make sure trigger cable from TCD plugged into Wavetek		____

8. Check PLC interlock                                                                       ____

# FEES & MWC FEES

1. Check that all blowers in rack row 2B are running		____

2. Turn on VME crate and boot processor				____

3. Make sure all manual switches on LVPS are in remote position	____

4. Turn on all FEEs and MWC FEEs	remotely.			____

5. Turn off TPC water skid – do all FEEs go off?			____

6. TPCTEMP computer running and updating?			____

# LASER

1. Power on both lasers – locally					____

2. Beams aligned?							____

3. Photodiode signal ok to TCD and cathode TDC?			____

4. Check remote on/off of lasers (slow controls)			____

5. Check remote viewing of lasers					____

6. Test PLC interlock							____

# TPC & GLOBAL INTERLOCKS

1. Check inner field cage air blower – on & flow switch ok?	____

2. Wet, in turn, east and west TPC Tracetek – alarm & skid stops?	____

3. Platform methane sniffer ok? Contact CAS if not			____

# GAS SYSTEM

1. Close SV18 and do leak rate test with N2				____

2. Calibrate and test trips of Pioneer methane system		____

3. Calibrate and test trip of gap methane detector			____

4. Calibrate CAI methane analyzers						____

5. Check limits and test hardware & PC alarms.			____

6. Bake out dryer & purifier							____

7. Check two big compressors					____

8. Test all gas pipe clamps and tighten				____

# ELECTRONICS

1. Do a geometry run – compare to baseline				____

2.Do a pulser run – compare to baseline				____

3. Do a full pedestal run – check RMS offline		                	_____