STAR Documented Work Procedure

SOP-TOF-GAS-03-A

OPERATING THE STAR TOF GAS SYSTEM

WITH r134a-isobutane-SF6

Text Pages 1 through 14

Attachment(s) 1,2,3,4,5,6,7,8,9,10

______________________                                          ______________________

Mike Gaffney                                                                 Bill Christie

ES&H Technical Review                                               STAR Safety Coordinator

_______________________				  _____________________

Geary Eppley					               Ralph Brown

STAR TOF Activities Manager			   Designated QA Representative

Preparer(s): Geary Eppley

STAR SOP-TOF-GAS-03-A									REV A

Date Issued: February 24, 2003

Category A

6.10.1.3 Operating the STAR TOF Gas System

1.0	Purpose and Scope

The scope of this procedure is those operations that are necessary for running

the STAR Time-of-Flight (TOF) with a gas of (90% r134a - 5% isobutene – 5% SF6).

Operations in this procedure include the following:

1. Flowing gas through the TOF system.

2. Monitoring the gas flow.

In the operation described here, r134a, isobutene, and SF6 gas are routed from the

gas storage pad, mixed in the gas system in the correct proportion and

sent to the TOF. The gas flow in the TOF is typically 70 ccm at a pressure

of 2 mbar. Critical flows of the gas system are monitored by a dedicated computer program..

The primary purpose of this operation is to initiate and maintain a precise mix

of the gas in the TOF for the duration of the STAR operations.

2.0	Responsibilities - See Attachment 1

3.0	Prerequisites

3.1	Required training to operate the gas system in the manner described in 1.0 				above:

3.1.1	BNL General Employee Training (GET)(V-001)

3.1.2	BNL Compressed gas safety course (OSH026)

3.1.3	BNL Cryogen Safety course (OSH025)

3.1.4	TOF gas system training with a previously qualified gas system

operator listed in Attachment 2 of this procedure.

- Procedure

# Please read this document in conjunction with the attached STAR engineering

# document TOFr-000-001 A, reference 7.1.

# The STAR TOFr gas system is a 3-component gas system that circulates gas

# through the TOFr tray and vents the gas to the atmosphere. TOFr may be

# operated with three gas mixtures: freon, freon-isobutane, and

# freon-isobutane-SF6. The gas is mixed in the ratio 90 part freon to 5

# parts isobutane to 5 parts SF6.

# To initiate gas flow in the system, the valve on the gas bottle should

# be fully opened and the regulator, MV 1, 2, or 3,  opened to produce about

# 14 pounds pressure in Pl 1, 2, or 3.

# The maximum flow for FM1 which controls the freon is 450 ccm. The maximum

flow for FM2, isobutane, and FM3, SF6, is 25 ccm. The freon flow should be

set on the control unit to the desired freon flow, 63 ccm, except when

initially charging the system. To introduce isobutane or SF6 into the

system, open the gas bottle and regulate the input pressure to

about 14 pounds. The control for FM2 (isobutane) or FM3 (SF6) should be

fully opened. Since FM2 and FM3 are slaved to FM1, FM2 and FM3 should flow

gas at the fixed ratio of 1:18 to the FM1 flow. For the normal freon flow

of 63 ccm, this is 3.5 ccm.

To prevent a failure of the flow meters from allowing a flammable gas

mixture to be introduced into STAR, the gas flow is monitored by a

computer which relays the information to slow controls once per

minute. The details of the interface may be found in the attached

document: TOF Systems in Slow Controls, attachment 3. Slow Controls will issue a red

alarm if the isobutane flow exceed 8% of the mixture and the shift crew

will be instructed to shut off the isobutane flow manually.

6.0	Documentation

6.1	All notes and observations should be recorded in the TOF gas system

logbook. A gas system log sheet (attachment 2) should be completed

every 8 hours and placed in the log sheet binder.

7.0	References

7.1	TOF mixing room schematic - STAR Drawing # TOFr-000-001 A

8.0	Attachments

1.	Responsible People/Operators, C-A-OPM-ATT 11.4.1.b

2.	Data Log Sheet

3.	Computer Alarm Conditions

5.